

PUMPWATCH CFX20/99

	
	

	PUMPWATCH PRODUCT SPECIFICATION
	PPS C209-2PK-L05-01

	
	
	

	CFX 99 Series 20L Mk.2 Single Measure Pack
	

	
	
	

	Product ref. C209-2PK
	

	
	
	

	General Description

	The PUMPWATCH 20L Single Measure Pack; component ref: C209-2PK comprises: 1 only 20L PUMPWATCH CFX 99 Series Mk.2 “Integrated Measure” type, volumetric Test Measure that when deployed, is suspended by means of a removable vertical tubular steel H-frame, from a 20L 2-wheel tubular steel Transport Trolley.  A 25mm diam. cam-lock type quick coupling Discharge System is provided to allow the contents of the Measure Pack to be discharged and drained after each test.
 
The Discharge System comprises a transparent, flexible, wire-reinforced polyurethane (pu) Discharge Hose that connects to the discharge valve of the 20L test measure and, is fitted on the discharge end with a male quick coupling connector that allows an additional Discharge Hose Extension to be connected if required.

	
	

	1.1
	1 x CFX 99 Series Mk.2 “integrated measure” type test measure each comprising a pre-preg carbon fibre composite construction main measure vessel otherwise the “main vessel”, and a pre-preg electro glass fibre composite construction measurement tube vessel, otherwise the “measurement tube”, bolted to the main vessel and connected to the upper neck of the main vessel by means of short transparent tube the “transfer tube” fitted with a 12mm ¼ turn ball-valve.  Each test measure is an instrument and is therefore delivered in a hinged, formed aluminium “transport case”. It is recommended, that the test measure be placed for security in the appropriate transport case whenever it is necessary to remove it from the transport trolley for storage, for calibration or for transport to or from the test site.

	1.1.1
	CFX 20/99 Mk.2 test measure – to contain 20L; comp. ref. C209-21
main vessel ID ø300mm;  measurement tube ID 38.1mm;  measurement tube scale length 350.8mm 
measurement tube scale resolution;  
-1.00% to 0 to +1.00% x 0.01%;  -200ml to 0 to +200ml x 2ml increments;
typical calibration accuracy: ??? limits of error ???on calibration ±0.02%
test measure overall height, ground level to top of funnel rim 1095mm;

overall height, base of discharge valve to top of funnel rim 790 ± 20; net wt. 3.8kg; 

gross wt incl. alu. transport case 11.4kg
transport case dimensions (external); L 880; H 365; W 380; wt. 7.6 kg

	1.2
	“20L” 2-wheel transport trolley; comp. ref. C209-T2 comprising:

	1.2.1
	2-wheel transport trolley chassis; comp.ref CM39-T2
 
comprising light-weight, green powder-coated tubular mild steel construction with 2 x vertical tubular shoes into which the removable H-frame legs are inserted; 2x 150mm diam x 50mm wide, solid, pneumatic wheels; U-profile tubular steel telescoping handle; adjustable height support legs;

outline dimensions: length (storage) 1,040mm; width (including wheels) 600mm; height without H-frame ; weight 9.6kg 

	
	“20L” 20L H-frame; comp. ref. C209-2202 comprising:

	1.2.2
	H-frame for suspending individual test measures from transport trolley chassis;
the removable H-frame comprising green powder-coated tubular mild steel construction shaped in large format H size width 352mm;660mm height; weight 1.200kg so called H-frame made of  2 x vertical tubular steel construction c/w with horizontal tubular to secure the test measure in level position of the vertical axis of the suspended test measure with butterfly level adjustment screw with plastic end cap and suspension leg end caps which the leg are inserted in the trolley shoes. 

	1.2.2
	H-frame x 1 for suspending 1 x test measures from the transport trolley chassis;
 
powder coated (green) tubular steel construction c/w 1 x horizontal M10 level adjustment screw, for independent adjustment of the position of the vertical axis of the 1 suspended test measures relative to the H-frame, so that the vertical axis of each test measure is precisely vertical when the test is being conducted, as follows;

	1.3
	Discharge System comp. ref C209-T3 
comprising transparent, flexible, wire reinforced polyurethane (pu) hoses with quick-coupling aluminium (alu) cam-lock type connectors as follows: 

	
	25mm nom diam x 1. m pu/aluminium discharge hose for discharging 20L test measure; c/w 135° elbow connection to 25mm quick-coupling female cam-lock type connector at Dispense End of hose, that connects to the discharge valve of the 20L test measure and; 25mm quick-coupling male cam-lock connector at Discharge End of hose; comp.ref. CM39-3301

	1.4
	Spare components (wear & tear) pack;  comp. ref. C209-29 comprising:

	1.4.1
	1x 20L/10L calibration piston o-ring (viton) 32.99mm IDxø2.62 cross section  comp.ref. CM39-110503 

	1.4.2
	2 off plastic end-cap ø10 x22 for level adjustmen screw; comp.ref. CM00-120205

	1.4.3
	2 off  M4 self-locking nut zn.plt DIN985 comp.ref. CM00-110002

	1.4.4
	2 off  A4 washer zn.plt ø4.5x12x1 comp.ref. CM00-110004

	
	

	
	

	2.
	Optional Items:

	2.1
	Calibration Stand for use when calibrating a test measure or checking the calibration of an already calibrated test measure comprising; 
Calibration Stand Base, a horizontal red powder-coated tubular steel base fitted with vertical “shoes” into which the vertical legs of a H-frame are inserted when assembling the stand;
Calibration Stand H-frame, a vertical red powder-coated tubular steel H-frame from which the test measure is suspended during calibration and c/w horizontal level adjustment screw for adjusting the position of the vertical axis of the suspended test measure relative to the H-frame, so that the axis is precisely vertical when the calibration is being conducted, as follows; 

	2.1.1
	20L/10L  2 Component Calibration Stand Assembly comp.ref. CM39-36 for use when calibrating 20L and 10L test measures; comprising; 

	2.1.2
	20L/10L Calibration Stand Base comp.ref. CM39-3601 
dimensions length 420 x 360 wide x 170 height; weight 3.36kg

	2.1.3
	20L/10L  Calibration Stand H-frame comp.ref. CM39-3602
dimensions   NOTE insert here and above “overall” and “between legs” dimensions 

	
	

	2.2
	Vapour Control Cap polyethylene (pe) construction to fit over the Filling Funnel of the test measure to prevent the escape of residual fuel fumes when the Measure Pack is not in use. Each cap comprises a formed pe lid with turn-down rim fitted internally with a bonded, flexible rubber foam sealing gasket. 

	2.2.1
	1 x 20L/10L Vapour Control Caps; comp.ref. CM39-2112 for 20L/10L test measures; funnel OD 240mm

	
	

	2.3
	Data Plate(s) Set as prescribed by DTI / NMO to be fixed to upper neck  of each test measure 
details of the Supplier and Client indicating model ref. nominal Volume at  temperature, and serial Number of the test measure  

	2.3.1
	Description aluminium Data Plate for attachment by means of 2 x pop rivets to upper neck of PUMPWATCH test measure. Gold painted, brushed aluminium with black printing and
2 x applications of clear lacquer to ensure that print is not removed by fuel typical dims.78 mm long x 
33 mm high  x  0.5mm thick

	2.4
	Discharge Hose Extension ø25x1.5m comp. ref. CM39-23

	2.4.1
	Description Ø25mmx1m with 135ºelbow and q/c connector

	
	

	
	

	
	


PPS C209-SPK –L05-01 MPG-AJB

